## Slide 1
Tensiometer Design Derren Lyons, Bailey McElhinney Department of Electrical & Computer Engineering, University at Albany
Problem Statement
The University at Albany’s chemistry department, as part of their educational mission, intends to build a laboratory for undergraduate students where they can gain hands on experience testing various material properties (including tensile strength). However, due to the extremely high cost of commercial products they have been unable to equip such a laboratory.
The goal of this project is to build an accurate, low cost, easy to use uniaxial tensile strength tester to support this educational mission.
System Requirements
Experimental Results
Project Partners
System Design
System Components & Budget
Metric | Sample | Our System | Known Value | Error
Young’s Modulus | Latex | 730 ± 10 kPa | 740 ± 10 kPa | 1.35%
… | Nitrile | 9.5 ± 0.1 MPa | 2.4 ± 0.2 MPa | 295.8%
Ult. Tensile Strength | Latex | 6.1 ± 0.1 MPa | 3.3 ± 0.1 MPa | 84.84%
… | Nitrile | 12.7 ± 0.1 Mpa | 4.4 ± 0.1 MPa | 188.63%
Part | Purpose | Cost
US-100 Ultrasonic Range Sensor | Measures Distance | $7
Capacity Hand Winch | Applies Force | $25
I2C 1602 LCD Display Module | Displays current measurement data | $10
TOTAL | TOTAL | $42
System Accuracy
Stress-Strain Curve
System Accuracy: The system should be able to accurately measure the stress-strain curve, such that Young’s Modulus and Ultimate Tensile Strength can be calculated to within two significant figures.
System Ease-of-Use: The system should be easy to calibrate and operate by undergraduate students without requiring knowledge of electronics or software programming.
Key System Features
To satisfy system requirements, we incorporated the following design modifications:
Item #1: Heavy duty hand winch with steel wire rope
Item #2: Various software improvements on the original code
Item #3: Improved ultrasonic range sensor
Item #4: LCD display module
Physical Model
Circuit Schematic
New and improved physical design model. Uniaxial application of force alongside real-time measurement data of force and distance displayed on display module.
Experimental results after implementation
of our system design.
Stress-Strain curve of a Nitrile sample obtained from our Arduino-based tester versus a commercial grade tester.
Circuit schematic of I2C LCD display module electronics (4 wires)
Stress-Strain curve of a Latex sample obtained from our Arduino-based tester versus a commercial grade tester.
Special thanks for Professor Chen, Feldblyun, Yeung, and the University at Albany’s Chemistry Department for sponsoring this project.
This project was developed in ECE442: Systems Analysis & Design in the Electrical & Computer Engineering Department.
Stress-Strain Curve